SPECIFICATION OF THE LASER MACHINE REQUIREMENT

Dual wavelength laser unit

Technical Specifications

  - Dual wavelength laser unit for treatment of varicose vein, hemorrhoid etc

  - Machine should emit dual wavelength laser of 980nm and 1470nm

  - Separately adjustable power output of 45 Watt (1470 nm / 15 Watt + 980 nm / 30 Watt)

  - Aiming beam: 532 nm and 635 nm | Green 1mW, Red 4nW, user controlled intensity

  - Should have multiple treatment modes which include but not limited to Continuous wave (CW), Pulse Mode and Derma mode)

  - Pulse duration /-break: 0.01 ‒ 60 sec / 0.01 ‒ 60 sec

  - User Interface: Touchscreen display for intuitive parameter control and preset procedures.

  - Cooling System: Air-cooled system for efficient operation.

  - Power supply: 110 – 240 VAC, 50 / 60 Hz, 450 VA

  - Should be light weight for easy carriage: ideally below 10 Kg

  - Should be US FDA approved or CE marked

  - Training: Training to doctors, anesthesiologists, biomedical engineers and nurses must be provided

Tumescent Infiltration Pump

Technical Specifications

- Volumetric displacement

  - Adjustable from 0–230ml/min. providing precise control for infiltration needs.

- Maximum pressure with closed tube set

  - Up to 2 bar

- Power Supply:

  - 100–240 V AC, 50–60 Hz,

- Motor:

  - High-performance, low-noise motor for consistent and reliable operation during long procedures.

- Control:

  - Intuitive control panel with easy-to-read display for setting and monitoring flow rate and pressure.

- Operation Mode:

  - Includes both manual adjustment and foot pedal for hands-free control.

- Tubing Compatibility:

  - Compatible with standard infiltration tubing and cannulas used in tumescent procedures.

- Safety Features:

  - Overpressure protection to prevent excessive delivery force.

  - Automatic shutoff in case of malfunction or overpressure detection.

- Material:

  - High-quality, medical-grade materials for durability and easy sterilization.

- Safety Standards:

  - Certified to comply with international medical device standards, including CE marking and FDA approval.

- Training

  - Training to doctors, anesthesiologists, biomedical engineers and nurses must be provided

Robotic Smart Pullback Device Unit

to automatically pullout fiber during procedure, for efficient outcome.

Technical Specifications

- Pulling speed

Adjustable speed of 0.5; 1.0; 1.5; 2.0 mm/s

- Fibers

  - ELVeS Radial®, or equivalent

  - ELVeS Radial 2ring® or equivalent

  - ELVeS Radial® slim, or equivalent

  - ELVeS Radial® 2ring slim or equivalent

- Power Supply

  - 100‒240 VAC, 50‒60 Hz

- Safety Standards:

  - Certified to comply with international medical device standards, including CE marking or US FDA approval.

- Training

  - Training to doctors, anesthesiologists, biomedical engineers and nurses must be provided

USG Scan Machine

- Portable Scan Machine with Trolley

- Probes

  - Vascular Probe

  - Soft Tissue Probe

- Training

  - Application Training for Users

  - Support Training for Biomedical Engineers

Consumable Requirements

Accessories and Consumables for the following Procedures

|  | Details | No of Procedures for 3 months |
| --- | --- | --- |
| 1 | Haemorrhoidectomy | 60 cases |
| 2 | Ana Fistula | 30 Cases |
| 3 | Pilonidal Sinus | 30 Cases |
| 4 | Varicose Vein | 20 Cases |

Format for rates of the consumables and accessories required for the procedures

|  | Item Code No | Detail of the item | Qty | Rate |
| --- | --- | --- | --- | --- |
| Haemorrhoidectomy |  |  |  |  |
|  |  |  | 1 |  |
|  |  |  | 1 |  |
|  |  |  | 1 |  |
| Ana Fistula |  |  |  |  |
|  |  |  | 1 |  |
|  |  |  | 1 |  |
| Pilonidal Sinus |  |  |  |  |
|  |  |  | 1 |  |
|  |  |  | 1 |  |
| Varicose Vein |  |  |  |  |
|  |  |  | 1 |  |
|  |  |  | 1 |  |