Indira Gandhi Memorial Hospital

Male’, Republic of Maldives

Biomedical Engineering Department

Technical Specification for Equipment

Technical Specification for Tabletop Pulse Oximeter

1. General Requirements:

- The Pulse Oximeter must be FDA approved and comply with international hospital-grade safety and performance standards.

- The device should be designed for continuous monitoring of oxygen saturation (SpO₂) and pulse rate for adult, paediatric, and neonatal patients.

- The equipment should have a shock-proof design to withstand accidental drops and impacts in a hospital environment.

- The system should be easy to operate, with an intuitive user interface and touchscreen or button controls.

2. Technical Specifications:

- SpO₂ Measurement Range: 0% – 100%

- SpO₂ Accuracy: ±2% (70% – 100%), unspecified below 70%

- Pulse Rate Measurement Range: 20 – 250 bpm

- Pulse Rate Accuracy: ±2 bpm or ±2%, whichever is greater

- Display: High-resolution LED or LCD with waveform (plethysmograph) and numerical values

- Alarms: Audible and visual alarms for high/low SpO₂ and pulse rate, sensor disconnection, and low battery

- Data Storage: Minimum 72 hours of trend data storage with recall feature

- Battery Backup: Minimum 4 hours of continuous operation on a rechargeable battery

- Power Supply: Compatible with 100-240V AC, 50/60Hz

- Connectivity: USB or wireless capability for data transfer and integration with hospital information systems

- Sensor Compatibility: Should support reusable and disposable SpO₂ sensors for adult, paediatric, and neonatal patients

3. Environmental Requirements:

- Operating Temperature: 5°C – 40°C

- Storage Temperature: -20°C – 60°C

- Humidity: 15% – 95% non-condensing

- Shock Resistance: Designed to withstand minor falls and impacts in a hospital setting

4. Accessories:

- Must include the following:

  - One (1) adult reusable SpO₂ sensor

  - One (1) paediatric reusable SpO₂ sensor

  - One (1) neonatal reusable SpO₂ sensor

  - One (1) disposable sensor pack (minimum 10 pieces)

  - One (1) power adapter and cable

  - One (1) battery pack

  - User manual and quick-start guide

5. Vendor Responsibilities:

- Must provide a one-year warranty covering parts and labour.

- Must provide installation and commissioning with an installation report upon completion.

- Must provide technical and user training to hospital staff.

- Must supply one (1) demonstration unit for trial use for a minimum period of seven (7) days.

- Must ensure after-sales support and service availability, including spare parts.

6. Compliance & Certifications:

- FDA Approved

- IEC 60601-1 (Medical Electrical Equipment - Safety)

- ISO 80601-2-61 (Particular requirements for Pulse Oximeters)

- CE Certification (if applicable)

End of Specification