Indira Gandhi Memorial Hospital

Male’, Republic of Maldives

Biomedical Engineering Department

Technical Specification for Equipment

Device Name: Digital Ear Thermometer

Intended Use: Measure human body temperature via the ear canal.

General Specifications

- Type: Infrared Digital Ear Thermometer

- Application: Hospital-grade, suitable for clinical and professional use

- Approval & Compliance: FDA-approved, IEC 60601-1 (Medical Electrical Equipment), ISO 80601-2-56 (Clinical Thermometry), CE Marked

- Measurement Method: Tympanic membrane infrared temperature detection

- Display Type: Large LCD with backlight

- User Interface: Single-button operation for ease of use

Measurement Performance

- Measurement Range: 32.0°C – 42.9°C (89.6°F – 109.2°F)

- Accuracy: ±0.2°C (±0.36°F) within 35.5°C – 42.0°C range

- Measurement Time: ≤ 2 seconds

- Measurement Mode:

  - Normal Mode (Quick single reading)

  - Continuous Mode (Multiple readings within a short duration)

- Memory Capacity: Stores at least 20 previous readings

Design & Build

- Shock-Proof Rating: Drop-resistant up to 1.5m (Tested as per IEC 60601-1 standards)

- Ergonomics: Lightweight and compact for easy handling

- Material: Medical-grade ABS with anti-microbial coating

- Probe Design:

  - Pre-calibrated, sensor for enhanced accuracy

  - Angled ergonomic probe for optimal ear canal fit

- Probe Cap Compatibility: Designed for universal disposable probe covers

- Protection: IP22 rating for protection against dripping water

Power & Battery

- Battery Type: 2 × AAA alkaline batteries

- Battery Life: At least 1,000 measurements per set

- Auto Power-Off: 30 seconds of inactivity

Safety & Compliance

- Non-invasive measurement ensuring patient safety

- Meets ASTM E1965-98(2016) standard for infrared ear thermometers

- Built-in fever alert (Audible & Visual Notification for High Temperature)

Additional Features

- Multi-User Mode: Suitable for different patient profiles

- C/F Conversion: Toggle between Celsius and Fahrenheit

- Hygienic Handling: Disposable probe covers with easy ejection system

Accessories

- Included in Package:

  - Digital Ear Thermometer Unit

  - Set of Disposable Probe Covers

  - Storage Case

  - User Manual

  - Calibration Certificate

Warranty: Minimum 6 months warranty covering manufacturing defects.

The vendor must provide a single sample unit/equipment as a demonstration unit for a minimum trial period of 7 days.

End of Specification